## Slide 1: lasFE
Lognormal Fitting of Laser Aerosol Spectrometer data for Filtration-Efficiency Analysis
Matt Wright
Research Associate I

## Slide 2: Laser Aerosol Spectrometer (LAS)
Optical spectrometer using intracavity laser to measure aerosol particulate size and number concentration [1]
User interface software has previously been approved for use by ICET’s Quality Assurance Group (QAG)
Test HEPA filters for radiological-waste treatment facilities
Data produced follows ICET internal procedures for measuring FE. This includes aerosol sampling upstream and downstream of the filter
GOAL: Analyze data to characterize filter’s ability to remove particulate from aerosol, i.e., Filtration Efficiency (FE)
GOAL: Minimize instrumental anomaly

## Slide 3: Computational Methods
Appropriately scale and plot Particle Size Distribution (PSD) by concentration and experimental conditions
Nonlinear Fit to Hind’s Frequency Function [2]:
  Count Mean Diameter (i.e., geometric mean, GM)
  Geometric Standard Deviation (GSD)
  Amplitude
Algorithms Attempted:
  Levenberg-Marquardt: minimization algorithm of an unbounded function [3]
  Dogbox: Trust Region algorithm optimized for large datasets relative to number of parameters [4]
  Arbitrary trial-and-error hard-coded parameters

## Slide 4: Fit Procedure
2-parameter fit for GM and GSD using scipy.optimize.curve_fit() [5]
Curve-area analysis to determine scaling
  Trapezoidal Integration used to determine area under PSD and fit [6]
  Fit scaled by area ratio such that PSD and fit functions have equivalent area
Both iterative methods fail to converge upon reasonable parameters. Going forward, considering 3-parameter fit to include scaling without curve-area analysis

## Slide 5: Levenberg-Marquardt (downstream dominant)

## Slide 6: Levenberg-Marquardt (downstream removed)

## Slide 7: Dogbox

## Slide 8: Good Fit

## Slide 9: Software Quality Assurance (SQA)
My inescapable Kafkaesque nightmare
Automated Calculation Application Packages (ACAP) – Excel data-reduction spreadsheets approved by QAG
  Limited iterative capabilities restricts complex analysis
  Blackbox GUI hides internal operations reducing operator control. May ultimately result in conditions adverse to quality
ICET-QA-036 Rev. 3, Software Control [7]

## Slide 10: SQA Activities Performed
Software Grade Level Determination (approved): A
  Graded approach ensures all levels of analysis are commensurate with safety, the DOE’s mission, or other relevant factors [8]
  Risk and Safety Assessment: Software failure may pose safety risks
Software Quality Assurance Plan (SQAP, in review)
  Outlines project-specific SQA practices and documentation
Software Requirements and Specifications (SRS, in preparation)
  Outlines functional and non-functional requirements and acceptance criteria
Test Plan and Execution (planned)
  Ensure acceptance criteria outlined in SRS is met

## Slide 11: Bibliography
TSI, “Laser Aerosol Spectrometer 3340A,”https://tsi.com/products/particle-sizers/supermicron-capable-particle-sizer-spectrometers/laser-aerosol-spectrometer-(las)-3340a/.
W. C. Hinds, Aerosol Technology, 2nd ed. (Wiley Inter-Science, 1999)
M. Lourakis, A Brief Description of the Levenberg-Marquardt Algorithm Implemented by Levmar 4 (2005)
K. Vogklis and I. Lagaris, “A rectangular trust-region approach for unconstrained and box-constrained optimization problems,” (2019) pp. 562–565.
Scipy Manual, optimize.curve_fit(), https://docs.scipy.org/doc/scipy/reference/generated/scipy. optimize.curve_fit.html#scipy.optimize.c
M. Newman, Computational Physics (2013)
ICET-QA-036 Rev. 3, Software Control
DOE O 414.1D Quality Assurance